# System 208V — Math formulas (typeset)

# System 208V — Math formulas

Typeset equations for R&D review (KaTeX). Companion files include the full Word mapping doc (System_208V_Column_Mapping_and_Math.docx) and the native-equation Word version (System_208V_Math_Formulas.docx). Source of truth remains the Rust pipeline under backend/src/processing/.

HTML + KaTeX · open in any browser

## Contents

- Shared math helpers

- Auto voltage modes

- Error % and phase Δ

- Load banding & averaging

- Comparison gradient

- Which format to use

## 1. Shared math helpers

Constants: 3=1.7320508075688772, near-zero ε=10−9.

### Three-phase average

avg(xA,xB,xC)=xA+xB+xC3

If any phase is missing → blank. Also used for Auto total THD percentages (reporting convention only).

### Unit scale (Auto power / reactive)

PkW=PW1000, SkVA=SVA1000, Qkvar=Qvar1000

### Reactive power (Auto)

Primary: instrument Q-* (signed), converted var → kvar.

Qkvar=Qvar1000

Fallback only if Q-* is blank/NAN and the triangle is valid:

Qfallback=S2−P2 when S2≥P2

Invalid triangle: if |P|>|S| beyond a tiny rounding tolerance, fallback is N/A — not |S2−P2|.

### Fallback total PF only

PFtotal=PFSIGMB/SIGMAif presentP/Sif PF blank and |S|>ε

Phase PFA/PFB/PFC use instrument fields only (no P/S fallback).

### Neutral proxy & unbalance (Auto)

IN=max(IA,IB,IC)−min(IA,IB,IC)

Uunbl(%)=maxi|xi−x||x|×100, x=avg(xA,xB,xC)

### Signed angle wrap

wrap(θ)∈(−180∘,180∘]

### Circular Δdeg

Δ=wrap​wrap(θmeter)−wrap(θauto)

## 2. Auto voltage modes

### IIR — line-to-neutral (ch 4/5/6)

UA&=Uac4, UB=Uac5, UC=Uac6ULN&=avg(Uac4,Uac5,Uac6)UAB&=Uac43, UBC=Uac53, UCA=Uac63ULL&=ULN3

Assumption: ideal balanced ULL=3 ULN so L–L columns can sit next to Acuvim L–L fields.

### IIW — line-to-line (ch 1/2/3)

UA,UB,UC,ULN&=N/AUAB&=Uac1, UBC=Uac2, UCA=Uac3ULL&=avg(Uac1,Uac2,Uac3)

### Phase displacement (Acuvim → Auto Phi)

ϕA,meter&=wrap(IAUA−UA)ϕB,meter&=wrap(IBUA−UB)ϕC,meter&=wrap(ICUA−UC)ϕA,auto&=wrap(−Phi4) (similarly for B/C / IIW)

If a phase voltage is missing, the corresponding current displacement is cleared (N/A), not left as a raw UA-relative angle.

## 3. Error % and phase averages

### Error percent

Error%=xmeter−xautoxauto×100

- Either side missing → N/A

- |xauto|≤10−9 → N/A (never invent zero)

Sign of Auto matters. When Auto is positive, positive Error% means the meter is algebraically higher. When Auto is negative (e.g. signed Q): meter =−9, Auto =−10:

Error%=−9−(−10)−10×100=−10%

Always read the formula, not only “higher / lower” wording.

### Band average — arithmetic (Real-Time / THD)

xc=1|Uc|i∈Uc​xi,c

### Band average — circular (Phase tables)

θ=atan2​1ni​sinθi, 1ni​cosθi then wrap to (−180∘,180∘]

Avoids the linear-mean bug: 179∘ and −179∘ average near ±180∘, not 0∘.

## 4. Load banding

Reference current: Auto SIGMB I after transform. For each sample and target Iktgt:

ek=|I−Iktgt|Iktgt×100

Keep targets with ek≤T (default T=5%). Assign to the smallest ek. Meter rows match by timestamp within 60 s default.

### Auto total THD (reporting convention)

UTHD,total=avg(UTHD,A,UTHD,B,UTHD,C)

R&D decision: this is the arithmetic mean of three phase THD percentages, not a physically aggregated THD of a combined waveform unless fundamentals are equal or the instrument defines “total” the same way.

## 5. Comparison cell gradient

Error % and Δdeg rows use a continuous green → yellow → red fill on absolute magnitude (Excel-style palette).

| Metric | Green | Yellow | Full red |
| --- | --- | --- | --- |
| Error % | ¦e¦=0%​ | ¦e¦=0.5%​ | ¦e¦≥1%​ |
| Δdeg | ¦Δ¦=0∘​ | ¦Δ¦=1.5∘​ | ¦Δ¦≥3∘​ |

Map |x| into t∈[0,1] with mid/high stops, then RGB-lerp green→yellow→red. See error_gradient_rgb in excel_write.rs.

## 6. Which document format to use

| Format | Best for | Equations |
| --- | --- | --- |
| This HTML | R&D math review, screenshots, print-to-PDF from browser | Typeset (KaTeX) |
| System_208V_Math_Formulas.docx | Word review with editable, native equation objects | Native Word OMML (generated by Pandoc) |
| System_208V_Column_Mapping_and_Math.docx | Email / Teams / Word comments, full column maps | Plain monospace (always readable) |
| COLUMN_MAPPING.md | Git history, code review | LaTeX source; only some viewers render it |

Native Word equations: Pandoc converts the LaTeX in this HTML to OMML equation objects. The generated Word file uses System_208V_Column_Mapping_and_Math.docx as its reference document so the page setup and document styles remain aligned with the full mapping guide.

Offline tip: first open needs CDN access for KaTeX; after cache, usually fine. Or File → Print → “Microsoft Print to PDF” for a portable PDF with equations already drawn.

RnD Data Processing · System 208V · formulas aligned with backend/src/processing