System 208V Accuracy Pipeline

Column mapping, units, and math (R&D review)

Purpose: Explain exactly how report values are produced so R&D can verify channel mapping, units, and formulas. Yokogawa Auto is the reference; Acuvim meters are the devices under test (DUT).

Formulas in this document are plain text (not LaTeX) so they display correctly in Word, Teams, and email.

# 1. Pipeline overview

Setup schedule (load % + target amps) + Data folder: Auto_*.CSV + Acuvim Real-Time (+ optional THD, PhaseAngle) ¦ v 1) Read Auto once; transform into meter-shaped columns per channel group 2) Segment load bands on Auto SIGMB current I(A) vs setup targets (+/- tolerance) 3) Match each meter (and THD/Phase) by timestamp; same load windows 4) Average 'used' samples per band (trim or window) 5) Compare: Error % or Delta deg 6) Write one Excel workbook per meter

# 2. Meter to Auto channel groups

| Acuvim meter | File pattern | Auto channels | Total | Voltage mode |
| --- | --- | --- | --- | --- |
| IIR / Meter 10 | *IIR*Real-Time*.csv | 4, 5, 6 | SIGMB | Line-to-neutral (L-N) |
| IIW / Meter 9 | *IIW*Real-Time*.csv | 1, 2, 3 | SIGMA | Line-to-line (L-L) |

| Report phase | IIR Auto channel | IIW Auto channel |
| --- | --- | --- |
| A | 4 | 1 |
| B | 5 | 2 |
| C | 6 | 3 |
| Total | SIGMB | SIGMA |

Load segmentation always uses Auto SIGMB (IIR-side current schedule), even for the IIW report. Both meters share the same time windows. Optional companions replace 'Real-Time' in the filename with 'THD' or 'PhaseAngle'. Exactly one Auto_*.CSV per folder.

# 3. Shared math helpers

Constants:  sqrt(3) = 1.7320508075688772    near-zero eps = 1e-9

## 3.1 Three-phase average

avg(xA, xB, xC) = (xA + xB + xC) / 3 If any phase is missing -> result is blank (N/A). Also used for Auto total U_THD / I_THD (reporting convention — see THD section).

## 3.2 Power unit conversion (Auto only)

P_kW = P_W / 1000 S_kVA = S_VA / 1000 Yokogawa P-* and S-* are in watts / VA; Acuvim Real-Time is already in kW / kVA.

## 3.3 Reactive power (Auto / WM side)

Primary (Yokogawa instrument Q): QA = Q-4 / 1000 QB = Q-5 / 1000 QC = Q-6 / 1000 (IIR) Q = Q-SIGMB / 1000 (IIR total) Same pattern for IIW with channels 1/2/3 + SIGMA. Units: Auto Q-* is in var; report uses kvar (divide by 1000). Sign is preserved (as Yokogawa reports lagging/leading). Fallback only if a Q-* cell is blank/NAN: Q = sqrt(S^2 - P^2) when S^2 >= P^2 (magnitude only) If ¦P¦ > ¦S¦ beyond tiny rounding tolerance -> N/A (invalid triangle) Total may also use sum of phase kvar if total Q-* is blank.

## 3.4 Fallback power factor

Total PF only: PF = PF-SIGMB (or PF-SIGMA), else P/S if blank (blank if ¦S¦ <= eps) Phase PFA/PFB/PFC: instrument PF-* only — no P/S fallback.

## 3.5 Neutral current proxy (Auto)

IN = max(IA, IB, IC) - min(IA, IB, IC) This is a spread proxy, not a true residual/neutral measurement.

## 3.6 Unbalance (Auto)

x_bar = avg(xA, xB, xC) Unbalance(%) = max(¦xi - x_bar¦) / ¦x_bar¦ * 100 Blank if ¦x_bar¦ <= eps. Used for U_UNBL(%) and I_UNBL(%) on WM Detail.

## 3.7 Signed angle wrap

wrap(theta) maps angle into (-180 deg, +180 deg] Example: wrap(355.9) = -4.1

## 3.8 Circular angle difference (Phase Comparison)

delta = wrap( wrap(theta_meter) - wrap(theta_auto) ) Result is always in (-180, +180] degrees.

## 3.9 Band average (per column)

Missing samples are skipped (not replaced with zero). If none remain -> N/A. Arithmetic mean (Real-Time, THD, non-phase): average_c = sum(values) / count Circular mean in degrees (Phase tables only — all columns are angles): mean = atan2( avg(sin theta), avg(cos theta) ) then wrap to (-180, 180] Avoids linear bug: 179 and -179 average to ~+/-180, not 0.

## 3.10 Error percent (Comparison and THD Comparison)

Error% = ( meter_avg - auto_avg ) / auto_avg * 100 Rules: - Either side missing -> N/A - ¦auto_avg¦ <= 1e-9 -> N/A (never invent 0 or silent zero fill) - When Auto is positive: positive Error% means meter is algebraically higher - When Auto is negative (e.g. signed Q): same formula, but everyday 'higher' can yield a negative Error%. Example: meter=-9, auto=-10 -> Error%=-10%

# 4. How Auto voltages are built

## 4.1 IIR — line-to-neutral mode (channels 4/5/6)

UA = Uac-4 UB = Uac-5 UC = Uac-6 ULN = avg(Uac-4, Uac-5, Uac-6) UAB = Uac-4 * sqrt(3) UBC = Uac-5 * sqrt(3) UCA = Uac-6 * sqrt(3) ULL = ULN * sqrt(3)

Review note: Assumes ideal balanced relation U_LL = sqrt(3) * U_LN so L-L columns can sit next to Acuvim L-L fields. If this wiring assumption is wrong for your setup, change the transform.

## 4.2 IIW — line-to-line mode (channels 1/2/3)

UA, UB, UC, ULN = N/A (not available in this mode) UAB = Uac-1 UBC = Uac-2 UCA = Uac-3 ULL = avg(Uac-1, Uac-2, Uac-3)

# 5. Real-Time column map

Acuvim Real-Time values are used as exported (no unit conversion). Auto is transformed into the same column names.

| Report column | Acuvim Real-Time | Auto IIR (4/5/6+SIGMB) | Auto IIW (1/2/3+SIGMA) |
| --- | --- | --- | --- |
| UA(V) | UA(V) as-is | Uac-4 | N/A |
| UB(V) | UB(V) | Uac-5 | N/A |
| UC(V) | UC(V) | Uac-6 | N/A |
| ULN(V) | ULN(V) | avg(Uac-4..6) | N/A |
| UAB(V) | UAB(V) | Uac-4 * sqrt(3) | Uac-1 |
| UBC(V) | UBC(V) | Uac-5 * sqrt(3) | Uac-2 |
| UCA(V) | UCA(V) | Uac-6 * sqrt(3) | Uac-3 |
| ULL(V) | ULL(V) | avg(Uac-4..6)*sqrt(3) | avg(Uac-1..3) |
| IA(A) | IA(A) | Iac-4 | Iac-1 |
| IB(A) | IB(A) | Iac-5 | Iac-2 |
| IC(A) | IC(A) | Iac-6 | Iac-3 |
| I(A) | I(A) | Iac-SIGMB (else avg I) | Iac-SIGMA (else avg I) |
| PA(kW) | PA(kW) | P-4 / 1000 | P-1 / 1000 |
| PB(kW) | PB(kW) | P-5 / 1000 | P-2 / 1000 |
| PC(kW) | PC(kW) | P-6 / 1000 | P-3 / 1000 |
| P(kW) | P(kW) | P-SIGMB / 1000 | P-SIGMA / 1000 |
| QA..Q (kvar) | as-is from meter | Q-4..6, Q-SIGMB / 1000 | Q-1..3, Q-SIGMA / 1000 |
| SA..S (kVA) | as-is | S-* / 1000 | S-* / 1000 |
| PFA..PFC | as-is | PF-4..6 only (no P/S fallback) | PF-1..3 only |
| PF total | as-is | PF-SIGMB else P/S | PF-SIGMA else P/S |
| FREQ(Hz) | FREQ(Hz) | FreqU-4 | FreqU-1 |
| IN(A) | as-is | maxI - minI | maxI - minI |
| U_UNBL / I_UNBL | as-is | unbalance formula | unbalance formula |

### Primary columns for pass/fail

- IIR: V (L-N), I, P, S, PF, F — treat Q Error% carefully (small denominator).

- IIW: V (L-L), I, total P/S, PF, F — phase UA.. often 0 on meter; phase PA.. Error% of -100% means meter exported 0, not a math bug.

# 6. THD column map

| Report column | Acuvim THD | Auto IIR | Auto IIW | App math |
| --- | --- | --- | --- | --- |
| UA_THD(%) | UA_THD(%) | Uthd-4 | Uthd-1 | as-is |
| UB_THD(%) | UB_THD(%) | Uthd-5 | Uthd-2 | as-is |
| UC_THD(%) | UC_THD(%) | Uthd-6 | Uthd-3 | as-is |
| U_THD(%) | U_THD(%) | avg Uthd 4..6 | avg Uthd 1..3 | Auto = average of 3 phases |
| IA_THD(%) | IA_THD(%) | Ithd-4 | Ithd-1 | as-is |
| IB_THD(%) | IB_THD(%) | Ithd-5 | Ithd-2 | as-is |
| IC_THD(%) | IC_THD(%) | Ithd-6 | Ithd-3 | as-is |
| I_THD(%) | I_THD(%) | avg Ithd 4..6 | avg Ithd 1..3 | Auto = average of 3 phases |

Not imported: odd/even THD, THFF, crest factor, K-factor (remain in raw THD CSV). Error% formula same as section 3.10.

Review note: Auto total U_THD / I_THD is the arithmetic mean of three phase THD percentages — a reporting convention, not a physically aggregated THD of a combined waveform unless fundamentals are equal or the instrument defines total the same way.

# 7. Phase column map and math

## 7.1 Why current angles are converted

Acuvim IB_UA / IC_UA are angles of IB/IC relative to voltage A, not relative to B/C. Yokogawa Phi-n is per-phase displacement (I vs that channel's U). The app converts before compare:

phi_A_meter = wrap( IA_UA - UA ) phi_B_meter = wrap( IB_UA - UB ) phi_C_meter = wrap( IC_UA - UC ) These still appear under headers IA_UA(deg), IB_UA(deg), IC_UA(deg) for layout continuity.

## 7.2 Auto Phi sign convention

For lagging near-unity PF on IIR, Acuvim displacement is typically about -4 deg while Yokogawa Phi is about +4 deg. Auto is stored as:

phi_auto_A = wrap( - Phi-4 ) // similarly Phi-5/6 or Phi-1/2/3 for IIW

Review note: The sign flip is convention alignment, not a claim that Yokogawa is wrong. If R&D defines lagging as positive, remove the negation and expect Delta deg near +/- 2*|phi| when signs disagree.

## 7.3 Phase Comparison third row: Delta deg

Delta_deg = circular_delta( meter_avg, auto_avg ) Not Error%. Cell fill uses the continuous gradient in section 10.1 (absolute deg scale).

## 7.4 Worked example (IIR full-load idea)

| Quantity | Value |
| --- | --- |
| Acuvim UA | 0 deg |
| Acuvim IA_UA | 355.9 deg |
| Auto Phi-4 | +4.03 deg |

phi_A_meter = wrap(355.9 - 0) = wrap(355.9) = -4.1 deg phi_A_auto = wrap(-4.03) = -4.03 deg Delta_deg ≈ -4.1 - (-4.03) = -0.07 deg

IIW Phase caution: on the dual-meter capture, Auto ch 1/2/3 phase powers can be highly unbalanced, so Phi-1..3 may look odd. Prefer IIR Phase + IIW Real-Time/THD for formal claims unless wiring validates IIW Phi.

# 8. Load banding, time match, averaging

## 8.1 Assign Auto rows to a load target

Reference: Auto SIGMB transformed I(A) (= Iac-SIGMB) For each sample current I and each setup target I_tgt: e = ¦I - I_tgt¦ / I_tgt * 100 Keep targets with e <= tolerance (default 5%). Assign sample to the target with the smallest e. If none within tolerance -> unmatched (detail sheet 'Unmatched rows').

## 8.2 Reduce: which samples inside a band are used

Trim mode: drop skip_start from the front and skip_end from the end; average the middle. If skips empty the band, use all samples. Window mode: skip skip_end from the end, then take the last window_size points before that. Short bands use what remains.

## 8.3 Timestamp match (meter / THD / Phase)

Default max ¦t_meter - t_auto¦ = 60 seconds (timestampMatchSeconds). Each meter sample joins the nearest Auto band timestamp if within 60 s. Then the same trim/window reduce is applied. THD and Phase reuse the same reference band times (no separate load detect).

# 9. Units summary

| Quantity | Acuvim CSV | Auto CSV | In report |
| --- | --- | --- | --- |
| Voltage | V | V | V |
| Current | A | A | A |
| Active power | kW | W | kW (= W/1000) |
| Apparent power | kVA | VA | kVA (= VA/1000) |
| Reactive | kvar (as-is) | var (Q-*) | kvar (= var/1000) |
| PF | unitless | unitless | unitless |
| Frequency | Hz | Hz (FreqU-*) | Hz |
| THD | % | % | % |
| Phase | deg | deg (Phi-*) | deg |

# 10. Excel sheets

| Sheets | Built from |
| --- | --- |
| Meter Detail, WM Detail, Comparison | Real-Time + Auto power transform |
| THD Meter / WM / Comparison | THD CSV + Auto Uthd/Ithd (if present) |
| Phase Meter / WM / Comparison | PhaseAngle CSV + Auto Phi (if present) |

Each Comparison load block: section header → WM AUTO average → METER average → Error% or Delta deg.

WM = working meter / Yokogawa reference.

## 10.1 Comparison Error % / Delta deg cell gradient

Error % and Delta deg rows use a continuous green → yellow → red fill on absolute magnitude (sign does not change color). Not three hard buckets.

Palette (Excel-style): green #63BE7B → yellow #FFEB84 → red #F8696B Error % scale: ¦e¦ = 0% green, 0.5% yellow, >= 1% full red Delta deg scale: ¦d¦ = 0 deg green, 1.5 deg yellow, >= 3 deg full red Map ¦x¦ into t in [0,1]: if x <= mid: t = 0.5 * (x / mid) if mid < x < high: t = 0.5 + 0.5 * ((x - mid) / (high - mid)) if x >= high: t = 1 t <= 0.5: RGB lerp green→yellow; t > 0.5: RGB lerp yellow→red N/A cells: neutral beige (no gradient). Code: excel_write.rs error_gradient_rgb

# 11. Talking points for R&D review

- Reference is always Yokogawa Auto; Error% = (meter - auto) / auto * 100.

- IIR = channels 4/5/6 + SIGMB; IIW = 1/2/3 + SIGMA.

- Load steps defined on SIGMB current; both meters time-aligned (±60 s default).

- Averages skip missing values (no silent zero fill).

- Auto P/S divided by 1000 to match Acuvim kW/kVA.

- IIR L-L Auto voltages = L-N * sqrt(3) (assumption — reviewable).

- Auto Q = Q-* / 1000 (var to kvar, signed); triangle only if Q blank and S^2>=P^2.

- Phase band averages use circular mean of degrees (not linear mean).

- Auto total THD = mean of three phase THD%s (reporting convention — confirm with R&D).

- THD is direct % compare; Auto totals = average of three phase THDs.

- Phase: meter converted to per-phase displacement; Auto Phi negated for lagging-sign align; report Delta deg.

- IIW phase P = 0 in Real-Time is export/wiring mode, not a division bug.

## Decisions R&D may want to change

| Decision | Current behavior | Possible alternative |
| --- | --- | --- |
| sqrt(3) L-L from L-N (IIR Auto) | Enabled | Blank L-L Auto; use L-N only |
| Auto Q source | Q-* / 1000 primary | Triangle fallback if blank |
| Auto Phi sign flip | Store -Phi | Keep +Phi; document lead/lag |
| Phase error metric | Delta deg circular | Also report ¦phi¦ Error% |
| Unbalance / IN on Auto | Derived proxies | Leave N/A if not instrument fields |
| Near-zero Auto denominator | N/A Error% | Flag band invalid |

# 12. Quick numeric spot-check

| Check | Expectation on a good full-load dwell |
| --- | --- |
| IIR I(A) Error% | typically much less than 1% |
| IIR P(kW) Error% | typically much less than 1% |
| IIR U_THD Error% | often under 1% |
| IIR Phase IA Delta deg | often under 1 deg after conversion + sign align |
| IIW total I / P Error% | typically << 1% mid loads; light loads can grow |

If full-load I/P Error% is about -99.9%, suspect a forgotten /1000 (should not happen in current code).

If Phase Delta deg is about 8 deg with |phi| about 4 deg, suspect sign convention not applied.

# 13. Code pointers

| Concern | Where to look |
| --- | --- |
| Channel groups | config/auto-channel-groups.json |
| Meter patterns / defaults | config/tests.registry.json |
| sqrt(3), P/S scale, Q, unbalance | backend/src/processing/preprocess.rs preprocess_auto_data |
| THD average | preprocess_auto_thd |
| Phase convert + Phi negate | preprocess_acuvim_phase, preprocess_auto_phase |
| Band assign | backend/src/processing/segment.rs segment_reference_bands |
| Time match | match_meter_bands |
| Average / Error% / Delta | backend/src/processing/compare.rs |
| Excel sheets | backend/src/processing/excel_write.rs |
| Comparison gradient fill | excel_write.rs → error_gradient_rgb |

When a report cell looks wrong: identify sheet → column → raw vs transformed Auto vs average vs Error/Delta → recompute with the equation in this document.